Repeat Cycle Prom

# Version aTl01WPCCfa
P/N ATL97DNTSTFA

## This EPROM will enable the user to run consecutive cycles

## This EPROM is fully compatible with all Tuttnauer EA and EZ units

Instructions for running repeat cycles:

- Turn the unit on and press START.

- The display will show Cyc Ctr: 000.

  - Using the UP/DOWN keys select the number of cycles to be run.

  - Then press the STOP key.

- The display will now show Dwell Tim: 000.

  - Using the UP/DOWN keys select the dwell time between cycles.

  - Press the STOP key.

  - The unit will now start the cycle.

Instructions for setting the automatic water fill:

- Press the STOP key repeatedly until the message “CODE:102” is displayed.

- Using the UP/DOWN keys advance to “CODE:105” and press the STOP key.

- The message “Water In: xxx Sec” will be displayed.

  - Using the UP/DOWN keys change the seconds according to the following:

  - 2340E/EK/EAEKA, EZ9 = 30 sec.

  - 2540E/EK/EAEKA, EZ10, EZ10k = 35 sec.

  - 3870E/EA = 65 sec.

  - After setting the correct time press the STOP key once.

- On some units the message “Ea Type:000” will appear.

  - Using the UP/DOWN select “0” for an E or EK type unit.

  - Using the UP/DOWN select “1” for an EA or EKA type unit.

  - Press the STOP key once to finish.

For more information, contact technical support at 1-800-762-1586

Ask For Burton Crawford or James Luo

Alfa Medical
265 Post Ave
Westbury, NY 11590